## 100K Ohm
| 742A-100K | PN010011 | 7630002 |  |  |  | Prediction Interval (95%) |  |  | ← PN010011 |  |  | 742A-100K | PN010010 | 7630001 |  |  |  | Prediction Interval (95%)
Date | Since Last Cal [Days] | Accumulated [Days] | Error [PPM] | Drifted [PPM] | Measured Value | MU (K =2) | Prediction | Lower | Upper | n | 21 | Date | Since Last Cal [Days] | Accumulated [Days] | Error [PPM] | Drifted [PPM] | Measured Value | MU (K =2) | Prediction | Lower | Upper
2012-08-06 | 0 | 0 | 4.87 |  | 100000.487 | 0.1 | 100000.4832 | 100000.4029 | 100000.5635 | df | 19 | 2017-05-09 | 0 | 0 | 5.2 |  | 100000.52 | 0.071 | 100000.5146 | 100000.4766 | 100000.5525
2013-02-13 | 191 | 191 | 5.5 | 0.63 | 100000.55 | 0.1 | 100000.4938 | 100000.4141 | 100000.5735 | Mean of X | 2352.6667 | 2017-12-27 | 232 | 232 | 5.5 | 0.3 | 100000.55 | 0.071 | 100000.5263 | 100000.4891 | 100000.5634
2013-09-06 | 205 | 396 | 4.8 | -0.7 | 100000.48 | 0.1 | 100000.5052 | 100000.4261 | 100000.5842 | Syx | 0.035585 | 2018-07-11 | 196 | 428 | 5.23 | -0.27 | 100000.523 | 0.046 | 100000.5362 | 100000.4996 | 100000.5727
2014-04-09 | 215 | 611 | 6.19 | 1.39 | 100000.619 | 0.12 | 100000.5171 | 100000.4386 | 100000.5956 | SSx | 47869864.6667 | 2019-02-19 | 223 | 651 | 5.48 | 0.25 | 100000.548 | 0.0046 | 100000.5474 | 100000.5113 | 100000.5834
2014-11-04 | 209 | 820 | 5 | -1.19 | 100000.5 | 0.12 | 100000.5287 | 100000.4507 | 100000.6067 |  |  | 2019-08-09 | 171 | 822 | 5.51 | 0.03 | 100000.551 | 0.0046 | 100000.556 | 100000.5203 | 100000.5917
2015-06-17 | 225 | 1045 | 5.2 | 0.2 | 100000.52 | 0.1 | 100000.5411 | 100000.4636 | 100000.6187 | PN010010 → |  | 2019-12-31 | 144 | 966 | 5.7 | 0.19 | 100000.57 | 0.0046 | 100000.5633 | 100000.5278 | 100000.5988
2016-01-12 | 209 | 1254 | 5.5 | 0.3 | 100000.55 | 0.1 | 100000.5527 | 100000.4756 | 100000.6299 | n | 14 | 2020-12-11 | 346 | 1312 | 5.8 | 0.1 | 100000.58 | 0.0046 | 100000.5807 | 100000.5455 | 100000.6159
2016-11-02 | 295 | 1549 | 5.2 | -0.3 | 100000.52 | 0.071 | 100000.5691 | 100000.4924 | 100000.6458 | df | 12 | 2021-09-01 | 264 | 1576 | 5.82 | 0.02 | 100000.582 | 0.0046 | 100000.594 | 100000.5588 | 100000.6292
2017-09-05 | 307 | 1856 | 5.4 | 0.2 | 100000.54 | 0.071 | 100000.5861 | 100000.5097 | 100000.6625 | Mean of X | 1439.6429 | 2022-04-07 | 218 | 1794 | 5.98 | 0.16 | 100000.598 | 0.0046 | 100000.605 | 100000.5696 | 100000.6404
2018-03-29 | 205 | 2061 | 5.9 | 0.5 | 100000.59 | 0.071 | 100000.5975 | 100000.5212 | 100000.6738 | Syx | 0.015609 | 2022-11-02 | 209 | 2003 | 6.1 | 0.12 | 100000.61 | 0.0046 | 100000.6155 | 100000.5799 | 100000.6512
2018-10-18 | 203 | 2264 | 5.92 | 0.02 | 100000.592 | 0.071 | 100000.6087 | 100000.5325 | 100000.685 | SSx | 11892727.2143 | 2023-06-20 | 230 | 2233 | 6.04 | -0.06 | 100000.604 | 0.0046 | 100000.6271 | 100000.5911 | 100000.6632
2019-05-13 | 207 | 2471 | 6.16 | 0.24 | 100000.616 | 0.0046 | 100000.6202 | 100000.544 | 100000.6965 |  |  | 2024-01-22 | 216 | 2449 | 6.22 | 0.18 | 100000.622 | 0.0046 | 100000.638 | 100000.6014 | 100000.6746
2020-05-19 | 372 | 2843 | 6.37 | 0.21 | 100000.637 | 0.0046 | 100000.6409 | 100000.5644 | 100000.7173 |  |  | 2024-10-07 | 259 | 2708 | 6.72 | 0.5 | 100000.672 | 0.0046 | 100000.6511 | 100000.6137 | 100000.6884
2021-05-06 | 352 | 3195 | 6.5 | 0.13 | 100000.65 | 0.0046 | 100000.6604 | 100000.5836 | 100000.7371 |  |  | 2025-07-07 | 273 | 2981 | 6.9 | 0.18 | 100000.69 | 0.0046 | 100000.6648 | 100000.6265 | 100000.7032
2021-12-20 | 228 | 3423 | 6.7 | 0.2 | 100000.67 | 0.0046 | 100000.673 | 100000.5959 | 100000.7501 |  |  | 2026-04-28 | 295 | 3276 | 6.87 | -0.03 | 100000.687 | 0.0046 | 100000.6797 | 100000.6401 | 100000.7193
2022-07-18 | 210 | 3633 | 6.71 | 0.01 | 100000.671 | 0.0046 | 100000.6847 | 100000.6072 | 100000.7621
2023-02-11 | 208 | 3841 | 7.18 | 0.47 | 100000.718 | 0.0046 | 100000.6962 | 100000.6183 | 100000.7741
2023-10-09 | 240 | 4081 | 6.87 | -0.31 | 100000.687 | 0.0046 | 100000.7095 | 100000.631 | 100000.788
2024-08-15 | 311 | 4392 | 7.5 | 0.63 | 100000.75 | 0.0046 | 100000.7268 | 100000.6474 | 100000.8061
2025-03-10 | 207 | 4599 | 7.81 | 0.31 | 100000.781 | 0.0046 | 100000.7382 | 100000.6583 | 100000.8182
2025-12-17 | 282 | 4881 | 7.59 | -0.22 | 100000.759 | 0.0046 | 100000.7539 | 100000.6729 | 100000.8348
2026-07-22 | 217 | 5098 | 7.79 | 0.2 | 100000.779 | 0.0022 | 100000.7659 | 100000.6841 | 100000.8477
Note on F6: [Threaded comment] Your version of Excel allows you to read this threaded comment; however, any edits to it will get removed if the file is opened in a newer version of Excel. Learn more: https://go.microsoft.com/fwlink/?linkid=870924 Comment: This data point observed to be outside the regression limits was measured years ago before the current analysis was implemented. Reply: @Siang-Yen Neng