## 10M Ohm
| 742A-10M | PN010015 | 7540002 |  |  |  | Prediction Interval (95%) |  |  | ← PN010015 |  |  | 742A-10M | PN010014 | 7510007 |  |  |  | Prediction Interval (95%)
Date | Since Last Cal [Days] | Accumulated [Days] | Error [PPM] | Drifted [PPM] | Measured Value | MU (K =2) | Prediction | Lower | Upper | n | 20 | Date | Since Last Cal [Days] | Accumulated [Days] | Error [PPM] | Drifted [PPM] | Measured Value | MU (K =2) | Prediction | Lower | Upper
2012-03-13 | 0 | 0 | 7.1 |  | 10000071 | 110 | 10000168.7626 | 10000072.7538 | 10000264.7715 | df | 18 | 2017-05-09 | 0 | 0 | 14.6 |  | 10000146 | 20 | 10000157.7872 | 10000116.6802 | 10000198.8943
2012-10-09 | 210 | 407 | 16.5 | 9.4 | 10000165 | 110 | 10000163.9376 | 10000069.4041 | 10000258.471 | Mean of X | 1949.15 | 2017-12-27 | 232 | 455 | 15.9 | 1.3 | 10000159 | 20 | 10000155.6536 | 10000115.9751 | 10000195.332
2013-05-01 | 0 | 0 | 23.2 | 6.7 | 10000232 | 110 | 10000168.7626 | 10000072.7538 | 10000264.7715 | Syx | 42.7401 | 2018-07-11 | 0 | 0 | 13.6 | -2.3 | 10000136 | 49 | 10000157.7872 | 10000116.6802 | 10000198.8943
2013-11-14 | 197 | 197 | 27.7 | 4.5 | 10000277 | 53.5 | 10000166.4272 | 10000071.1725 | 10000261.6818 | SSx | 40753700.55 | 2019-02-19 | 223 | 223 | 17.04 | 3.44 | 10000170.4 | 49 | 10000156.7415 | 10000116.4226 | 10000197.0604
2014-06-30 | 228 | 425 | 18.8 | -8.9 | 10000188 | 53.5 | 10000163.7242 | 10000069.2485 | 10000258.1998 |  |  | 2019-08-12 | 174 | 397 | 13.66 | -3.38 | 10000136.6 | 49 | 10000155.9256 | 10000116.1047 | 10000195.7465
2014-12-26 | 179 | 604 | 10.68 | -8.12 | 10000106.8 | 20 | 10000161.6021 | 10000067.6657 | 10000255.5384 | PN010014 → |  | 2020-01-02 | 143 | 540 | 17.65 | 3.99 | 10000176.5 | 49 | 10000155.255 | 10000115.7636 | 10000194.7464
2015-08-05 | 222 | 826 | 12.4 | 1.72 | 10000124 | 20 | 10000158.9702 | 10000065.6127 | 10000252.3277 | n | 14 | 2020-12-11 | 344 | 884 | 17.38 | -0.27 | 10000173.8 | 49 | 10000153.6419 | 10000114.6372 | 10000192.6466
2016-04-12 | 251 | 1077 | 14.6 | 2.2 | 10000146 | 20 | 10000155.9945 | 10000063.1692 | 10000248.8199 | df | 12 | 2021-09-02 | 265 | 1149 | 13.52 | -3.86 | 10000135.2 | 49 | 10000152.3992 | 10000113.4675 | 10000191.3309
2017-09-05 | 511 | 1588 | 12.2 | -2.4 | 10000122 | 20 | 10000149.9365 | 10000057.7852 | 10000242.0878 | Mean of X | 1089.4286 | 2022-04-08 | 218 | 1367 | 17.4 | 3.88 | 10000174 | 49 | 10000151.377 | 10000112.3066 | 10000190.4473
2018-03-29 | 205 | 1793 | 14.7 | 2.5 | 10000147 | 20 | 10000147.5062 | 10000055.4688 | 10000239.5436 | Syx | 17.2595 | 2022-11-03 | 209 | 1576 | 15.2 | -2.2 | 10000152 | 49 | 10000150.3969 | 10000111.027 | 10000189.7668
2018-10-18 | 203 | 1996 | 14.2 | -0.5 | 10000142 | 20 | 10000145.0996 | 10000053.0861 | 10000237.1131 | SSx | 9610883.4286 | 2023-06-20 | 229 | 1805 | 15.6 | 0.4 | 10000156 | 49 | 10000149.3231 | 10000109.442 | 10000189.2041
2019-05-13 | 207 | 2203 | 14.95 | 0.75 | 10000149.5 | 49 | 10000142.6455 | 10000050.5651 | 10000234.726 |  |  | 2024-01-24 | 218 | 2023 | 16.1 | 0.5 | 10000161 | 49 | 10000148.3008 | 10000107.762 | 10000188.8396
2020-05-19 | 372 | 2575 | 14.1 | -0.85 | 10000141 | 49 | 10000138.2354 | 10000045.8041 | 10000230.6667 |  |  | 2024-10-07 | 257 | 2280 | 13.3 | -2.8 | 10000133 | 49 | 10000147.0956 | 10000105.5779 | 10000188.6134
2021-05-07 | 353 | 2928 | 15.28 | 1.18 | 10000152.8 | 49 | 10000134.0505 | 10000041.0149 | 10000227.0861 |  |  | 2025-07-07 | 273 | 2553 | 12.8 | -0.5 | 10000128 | 49 | 10000145.8155 | 10000103.033 | 10000188.5979
2021-12-20 | 227 | 3155 | 11.6 | -3.68 | 10000116 | 49 | 10000131.3594 | 10000037.798 | 10000224.9208 |  |  | 2026-04-28 | 295 | 2848 | 13.4 | 0.6 | 10000134 | 49 | 10000144.4321 | 10000100.0452 | 10000188.819
2022-07-21 | 213 | 3368 | 11.8 | 0.2 | 10000118 | 49 | 10000128.8342 | 10000034.6836 | 10000222.9848
2023-02-11 | 205 | 3573 | 14.5 | 2.7 | 10000145 | 49 | 10000126.4039 | 10000031.6001 | 10000221.2076
2023-10-09 | 240 | 3813 | 13.9 | -0.6 | 10000139 | 49 | 10000123.5586 | 10000027.8854 | 10000219.2318
2024-08-15 | 311 | 4124 | 11.3 | -2.6 | 10000113 | 49 | 10000119.8716 | 10000022.9085 | 10000216.8348
2025-03-10 | 207 | 4331 | 11.8 | 0.5 | 10000118 | 49 | 10000117.4176 | 10000019.4969 | 10000215.3383
2025-12-17 | 282 | 4613 | 14.8 | 3 | 10000148 | 49 | 10000114.0744 | 10000014.7266 | 10000213.4222
Note on F6: [Threaded comment] Your version of Excel allows you to read this threaded comment; however, any edits to it will get removed if the file is opened in a newer version of Excel. Learn more: https://go.microsoft.com/fwlink/?linkid=870924 Comment: This data point observed to be outside the regression limits was measured years ago before the current analysis was implemented.